## sbRIO
sbRIO9637 Board Signals
J4 DIO pin | J5 MIO pin | Signal | Notes
2 DIO4 |  | AD 0 | ADC Bit 0 (LSB)
4 DIO5 |  | AD 1 | ADC Bit 15 (MSB)
6 DIO6 |  | AD 2 | ADC Bit 2
8 DIO7 |  | AD 3 | ADC Bit 3
10 DIO8 |  | AD 4 | ADC Bit 4
12 DIO9 |  | AD 5 | ADC Bit 5
14 DIO10 |  | AD 6 | ADC Bit 6
16 DIO11 |  | AD 7 | ADC Bit 7
18 DIO12 |  | AD 8 | ADC Bit 8
20 DIO13 |  | AD 9 | ADC Bit 9
22 DIO14 |  | AD 10 | ADC Bit 10
24 DIO15 |  | AD 11 | ADC Bit 11
26 DIO16 |  | AD 12 | ADC Bit 12
28 DIO17 |  | AD 13 | ADC Bit 13
30 DIO18 |  | AD 14 | ADC Bit 14
32 DIO19 |  | AD 15 | ADC Bit 15 (MSB)
34 DIO20 |  | AD OF | ADC over/under flow bit. High when OF/UF condition.
36 DIO21 |  | AD CLK+ | ADC clock + output (use to clock data into sbRIO9636)
38 DIO22 |  | AD CLK- | ADC clock - output (use to clock data into sbRIO9636)
40 DIO23 |  | AD EN | ADC enable (logic low to enable)
42 DIO24 |  | ENCODE | ADC encode input (25 MHz from 9683)
46 DIO26 |  | DIO_26 | Plenum Fan Bank 4
48 DIO27 |  | DIO_27 | Digital I/O reserved for expansion
1,3,5,…47 |  | DGND | Digital Grounds
49 5V |  | 5 V | 5 volts
50 5V |  | 5 V | 5 volts
| 2 AI0 | Temperature 1 | Thermistor 1 Voltage (0 to +10V)
| 6 AI1 | Temperature 2 | Thermistor 2 Voltage (0 to +10V)
| 8 AI2 | Temperature 3 | Thermistor 3 Voltage (0 to +10V)
| 12 AI3 | Temperature 4 | Thermistor 4 Voltage (0 to +10V)
| 14 AI4 | Sample Temp | Vaisala probe #1 (0 to +5V)
| 18 AI5 | Sample RH | Vaisala probe #1 (0 to +5V)
| 20 AI6 | External Temp | Vaisala probe #2 (0 to +5V)
| 24 AI7 | External RH | Vaisala probe #2 (0 to +5V)
| 3 AI8 |  | Analog Input reserved for expansion
| 5 AI9 |  | Analog Input reserved for expansion
| 9 AI10 |  | Analog Input reserved for expansion
| 11 AI11 |  | Analog Input reserved for expansion
| 15 AI12 |  | Analog Input reserved for expansion
| 17 AI13 |  | Analog Input reserved for expansion
| 21 AI14 |  | Analog Input reserved for expansion
| 23 AI15 |  | Analog Input reserved for expansion
| 28 AO0 | KNF pump | 0 -5 V @ x A
| 30 AO1 | primary volume fan | 0 - 5 V to V to PWM converter
| 32 AO2 | AO_2 | Analog Output reserved for expansion
| 34 AO3 | AO_3 | Analog Output reserved for expansion
| 44 DIO0 | Laser Mod | Laser Modulation (low to turn laser on)
| 46 DIO1 | Motor Control | PWM to motor
| 48 DIO2 | DIO_2 | Plenum Fan Bank 2
| 50 DIO3 | DIO_3 | Plenum Fan Bank 3
ASRL1 | W2 | RS232 | Buck Hygrometer
ASRL2 | W4 | RS232 | Alicat Flow Controller
ASRL3 | W3 | RS485 | reserved

## Power
AC power inventory |  | AC current | Notes
OPCRD Electronics |  | tbd
DC power inventory
Device | voltage | current
PMT | +5 V | 25 mA | Hamamatsu HM9433-201
| -5 V | 25 mA
Plenum fans | +5 V | .16 A x 9 = 1.44 A | NMB-MAT 1608KL-01W-B20 6.0 cfm x 9 = 54 cfm
Electronic Box fans | +5 V | .16 A x 2 = .32 A | NMB-MAT 1608KL-01W-B20 6.0 cfm each
Laser | +5 V
Servo | +5 V
sbRIO | +12 V | 1 A
temperature sensor #1 | +12 V |  | Precision plenum temperature
temperature sensor #2 | +12 V |  | Precision plenum temperature
temperature sensor #3 | +12 V |  | Coarse external temperature
temperature sensor #4 | +12 V
Vaisala probe #1 | +12 V
Vaisala probe #2 | +12 V
Alicat flow controller | +12 V
Rotronics HygroClip2 probe | +3.3 V | 4.5 mA | regulated from +12 V supply
Buck CR-5 Hygrometer | +24 V | 32A | TDK-Lambda SCS120PW12 5.5"x3"x1.5"
sample heater | +24 V | 2 A
enclosure fans (2x) | +24 V | .3 A | eg NMB 3610KL-05W-B50-D00 (54 cfm max)
Air flow meter | +24V | 50 mA | Flow Networks 3202.30
KNF NMP 015 B | +6 V | .3 A max | regulated from +24 V supply

## Panel connector
Designator | Description | Manufactrurer | Model
P1 | AC power recepticle | Shurter | 6100-3300
CB1 | circuit breaker | Klixon | tbd
J8 | Ethernet (RJ45) | L-Com | ECF504-SC5E
| USB | L-Com | ECF504-AA
D9-S RS232
D9-S RS232
D9-S RS485