## sbRIO
sbRIO9637 Board Signals
J4 DIO pin | J5 MIO pin | Signal | Notes
2 DIO4 |  | AD 0 | ADC Bit 0 (LSB)
4 DIO5 |  | AD 1 | ADC Bit 15 (MSB)
6 DIO6 |  | AD 2 | ADC Bit 2
8 DIO7 |  | AD 3 | ADC Bit 3
10 DIO8 |  | AD 4 | ADC Bit 4
12 DIO9 |  | AD 5 | ADC Bit 5
14 DIO10 |  | AD 6 | ADC Bit 6
16 DIO11 |  | AD 7 | ADC Bit 7
18 DIO12 |  | AD 8 | ADC Bit 8
20 DIO13 |  | AD 9 | ADC Bit 9
22 DIO14 |  | AD 10 | ADC Bit 10
24 DIO15 |  | AD 11 | ADC Bit 11
26 DIO16 |  | AD 12 | ADC Bit 12
28 DIO17 |  | AD 13 | ADC Bit 13
30 DIO18 |  | AD 14 | ADC Bit 14
32 DIO19 |  | AD 15 | ADC Bit 15 (MSB)
34 DIO20 |  | AD OF | ADC over/under flow bit. High when OF/UF condition.
36 DIO21 |  | AD CLK+ | ADC clock + output (use to clock data into sbRIO9636)
38 DIO22 |  | AD CLK- | ADC clock - output (use to clock data into sbRIO9636)
40 DIO23 |  | AD EN | ADC enable (logic low to enable)
42 DIO24 |  | ENCODE | ADC encode input (25 MHz from 9683)
44 DIO25 |  | DIO_25 | Digital I/O reserved for expansion
46 DIO26 |  | DIO_26 | Plenum Fan Bank 4
48 DIO27 |  | DIO_27 | Digital I/O reserved for expansion
1,3,5,…47 |  | DGND | Digital Grounds
49 5V |  | 5 V | 5 volts
50 5V |  | 5 V | 5 volts
| 2 AI0 | Temperature 1 | Thermistor 1 Voltage (0 to +5V)
| 6 AI1 | Temperature 2 | Thermistor 2 Voltage (0 to +5V)
| 8 AI2 | Temperature 3 | Thermistor 3 Voltage (0 to +5V)
| 12 AI3 | Temperature 4 | Thermistor 4 Voltage (0 to +5V)
| 14 AI4 | Plenum Temp | Vaisala probe #1 (0 to +5V)
| 18 AI5 | Plenum RH | Vaisala probe #1 (0 to +5V)
| 20 AI6 | Ambient Temp | Vaisala probe #2 (0 to +5V)
| 24 AI7 | Ambient RH | Vaisala probe #2 (0 to +5V)
| 3 AI8 | Air Flow | 4-20 mA FlowNetwork Airflow sensor 250 ohm resistor
| 5 AI9 | Pressure | Omega PX142-015A5V P = 2 (V-1.5) for 12V supply
| 9 AI10 |  | Analog Input reserved for expansion
| 11 AI11 |  | Analog Input reserved for expansion
| 15 AI12 |  | Analog Input reserved for expansion
| 17 AI13 |  | Analog Input reserved for expansion
| 21 AI14 |  | Analog Input reserved for expansion
| 23 AI15 | Sample heater temp | Sample Heater Thermocouple T = 200 (V-1.25)
| 28 AO0 | KNF pump | 0 -5 V @ x A
| 30 AO1 | Primary volume fan | 0 - 1 V to V to PWM converter
| 32 AO2 | Sample heater control | 0 - 1 V to V to PWM converter
| 34 AO3 | AO_3 | Analog Output reserved for expansion
| 44 DIO0 | Laser Mod | Laser Modulation (low to turn laser on)
| 46 DIO1 | Servo Control | PWM to servo
| 48 DIO2 | DIO_2 | Plenum Fan Bank 2
| 50 DIO3 | DIO_3 | Plenum Fan Bank 3
ASRL1 | W2 | RS232 | Buck Hygrometer
ASRL2 | W4 | RS232 | Alicat Flow Controller
ASRL3 | W3 | RS485 | reserved

## Power
AC power inventory |  | AC current | Notes
OPCRD Electronics |  | tbd | 2 A cicuit breaker Klixon 7277-2-2
DC power inventory
Device | voltage | current
PMT | +5 V/-5 V | 25 mA/25 Ma | Hamamatsu HM9433-201
Plenum fans | +5 V | .16 A x 9 = 1.44 A | NMB-MAT 1608KL-01W-B20 6.0 cfm x 9 = 54 cfm
Electronic Box fans | +5 V | .16 A x 2 = .32 A | NMB-MAT 1608KL-01W-B20 6.0 cfm each
electronic volume fan | +5 V | 120 mA | ADDA AD0605LB-D71GL 12 cfm
Laser | +5 V | 50 mA
Servo | +5 V
sbRIO | +12 V | 1 A
temperature sensor #1 | +12 V |  | Precision plenum temperature
temperature sensor #2 | +12 V |  | Precision plenum temperature
temperature sensor #3 | +12 V
temperature sensor #4 | +12 V
Vaisala probe #1 | +12 V | 5 mA
Vaisala probe #2 | +12 V | 5 mA
Alicat flow controller | +12 V | .25 A
Omega PX142-015A5V | +12 V | 20 mA
Buck CR-5 Hygrometer | +24 V | 2 A | TDK-Lambda SCS120PW12 5.5"x3"x1.5"
sample heater | +24 V | 2 A
primary volume fan | +24 V | .3 A | NMB 3610KL-05W-B50-D00 (54 cfm max)
Air flow meter | +24V | 50 mA | Flow Networks 3202.30
KNF NMP 015 B | +6 V | .3 A max | regulated from +24 V supply

## Panel connectors
Designator | Description | Manufactrurer | Part Number | Notes
J1 | AC power recepticle | Shurter | 6100-3300 | 1/4" tabs
J2 | Electronic Volume Fan | Switchcraft | EN2P2F20 | 2-pin mating connector EN2C2M20DC
J3 | Plenum Fans | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J4 | PMT Monitor | Amphenol | 31-220N | BNC
J5 | Primary Volume Fan | Switchcraft | EN2P2F20 | 2-pin mating connector EN2C2M20DC
J6 | Air Flow Meter | Switchcraft | EN2P3F20 | 3-pin mating connector EN2C3M20DC
J7 | Vaisalla Plenum T/RH Probe | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J8 | Temperature Probe 1 | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J9 | Temperature Probe 2 | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J10 | Laser Modulation | Amphenol | 31-220N | BNC
J11 | Laser Power | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J12 | Servo | Switchcraft | EN2P3F20 | 3-pin mating connector EN2C3M20DC
J13 | Sample Pump | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J14 | Sample Heater | Switchcraft | EN2P2F20 | 2-pin mating connector EN2C2M20DC
J15 | Plenum Flow Indicator LED | Switchcraft | EN2P2F20 | 2-pin mating connector EN2C2M20DC
J16 | Power Indicator LED | Switchcraft | EN2P2F20 | 2-pin mating connector EN2C2M20DC
J17 | Heater Thermocouple |  |  | 7/8" dia round K-type connector
J18 | Ethernet Connector | L-Comm | ECF504-SC5E | Shielded RJ-45
J19 | USB Connector | L-Comm | ECF504-AA
J20 | Flow Control | Multicomp | SPC15477 | D9-F
J21 | Vaisalla External T/RH Probe | Switchcraft | EN2P4F22 | 4-pin mating connector EN2C4M22DC
J22 | RS485 | Multicomp | SPC15477 | D9-F
CB1 | Circuit Breaker | Klixon | 7277-2-2 | 14500-5 water seal
| PMT Gain | Bourns | H-23-6A | Potentiometer Counting Dial