## Sheet1
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN | CN 74000 | CN
[8 empty] | 78000 |  | 79000
2 | 3 |  | && 00 |  | Process Value | Read Only | Actual Input Value | X | X | X
3 | 3 |  | && 00 |  | Unit Status | Read Only | Binary value, details below | X | X | X
[6 empty] | [High Byte] | bit7 Auto/Manual mode, 0=Auto 1=Man |  | X
[7 empty] | bit6 Comm. mode 0=Local 1=Remote | X | X | X
bit5 Not used
[7 empty] | bit4 1=Error present, read register 0x0003 | X | X | X
[7 empty] | bit3 Alarm 1 output 0=Off 1=On | X | X
[7 empty] | bit2 Alarm 2 output 0=Off 1=On |  | X
bit1, bit0 Local set point selected
0,0=1SP1 0,1=2SP1
1,0=3SP1 1,1=4SP1
[6 empty] | [Low Byte] | bit7 1=NAT error (no activity timer) | X | X | X
bit6,5,4 Decimal point for Process Value
[7 empty] | 000=0 001=0.0 010=0.00 011=0.000 | X | X | X
[7 empty] | 100=0.0000 (bit6 for PP45,16A only) | X
bit3 Not used
[7 empty] | bit2,bit1 Engineering units | X | X | X
00=None 01=Deg. F 10=Deg. C
[7 empty] | bit0 Input value sign 0= Pos. 1= Neg. | X | X | X
4 | 3 |  | && 05 |  | Unit Error Status | Read Only | Binary value, details below
[6 empty] | [High Byte] | bit7 Unit Failed Self-test | X | X | X
bit6 Not used
[7 empty] | bit5 Unit Calibration bad | X | X | X
[7 empty] | bit4 Input overflow | X | X | X
[7 empty] | bit3 Input underflow | X | X | X
[7 empty] | bit2 Bad input | X | X | X
[7 empty] | bit1 Open input | X | X | X
[7 empty] | bit0 Unit Ambient Temp. beyond spec. | X | X | X
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
[6 empty] | [Low Byte] | bit7 Control Loop break | X | X | X
[7 empty] | bit6 Sensor rate of change exceeds limit | X | X | X
bit5,bit4,bit3,bit2,bit1,bit0 Not used
7 | 3 |  | && 011F |  | Active Segment | Read Only | Programmer segment 1 to 16 |  | X
8 | 3 |  | && 0120 |  | TBAS | Read Only | 0 = 1 second 1 = 60 seconds |  | X
257 | 3 |  | 100 |  | Active Set Value | Read Only | **** See note | X | X | X
258 | 3 | 6 | 101 | 200 | SP1 or 1SP1 | R/W | **** See note | X | X | X
259 | 3 | 6 | 102 | 201 | 2SP1 | R/W | **** See note
260 | 3 | 6 | 103 | 202 | 3SP1 | R/W | **** See note
261 | 3 | 6 | 104 | 203 | 4SP1 | R/W | **** See note
262 | 3 | 6 | 105 | 204 | SP2 | R/W | **** See note | X | X
263 | 3 | 6 | 106 | 205 | A1LO | R/W | **** See note | X | X
264 | 3 | 6 | 107 | 206 | A1HI | R/W | **** See note | X | X
265 | 3 | 6 | 108 | 207 | A2LO | R/W | **** See note
266 | 3 | 6 | 109 | 208 | A2HI | R/W | **** See note
267 | 3 | 6 | 010A | 022B | SP1D | R/W |  | X | X | X
268 | 3 | 6 | 010B | 022C | SP2D | R/W |  | X | X
269 | 3 | 6 | 010C | 022F | PB1 or 1PB1 | R/W | !! See note | X | X | X
270 | 3 | 6 | 010D | 230 | 2PB1 | R/W | !! See note
271 | 3 | 6 | 010E | 231 | 3PB1 | R/W | !! See note
272 | 3 | 6 | 010F | 232 | 4PB1 | R/W | !! See note
273 | 3 | 6 | 110 | 023F | PB2 | R/W | !! See note | X | X
274 | 3 | 6 | 111 | 233 | RES or 1RES | R/W | - # = reset + # = offset | X | X | X
275 | 3 | 6 | 112 | 234 | 2RES | R/W | - # = reset + # = offset
276 | 3 | 6 | 113 | 235 | 3RES | R/W | - # = reset + # = offset
277 | 3 | 6 | 114 | 236 | 4RES | R/W | - # = reset + # = offset
278 | 3 | 6 | 115 | 237 | RTE or 1RTE | R/W | WR 1 for 00.01 minutes WR 0 = OFF | X | X | X
279 | 3 | 6 | 116 | 238 | 2RTE | R/W | WR 1 for 00.01 minutes WR 0 = OFF
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
280 | 3 | 6 | 117 | 239 | 3RTE | R/W | WR 1 for 00.01 minutes WR 0 = OFF
281 | 3 | 6 | 118 | 023A | 4RTE | R/W | WR 1 for 00.01 minutes WR 0 = OFF
282 | 3 | 6 | 119 | 240 | ARTE | R/W | WR 1 for 00.01 minutes WR 0 = OFF | X | X | X
283 | 3 | 6 | 011A | 241 | FINT | R/W | WR 1 for 1% WR 0 = OFF | X | X | X
284 | 3 | 6 | 011B | 242 | FBND | R/W |  | X | X | X
285 | 3 | 6 | 011C | 243 | FRTE | R/W | WR 1 for 00.01 counts/second | X | X | X
286 | 3 |  | 011D |  | PEA | Read Only | !! See note | X | X | X
287 | 3 |  | 011E |  | VAL | Read Only | !! See note | X | X | X
288 | 3 |  | && 011F |  | Ti | Read Only | Time remaining in segment. |  | X
289 | 3 | 6 | 120 | 209 | 1Ti | R/W | Time for segment. |  | X
290 | 3 | 6 | 121 | 219 | 1SP | R/W | **** See note |  | X
291 | 3 | 6 | 122 | 020A | 2Ti | R/W | Time for segment. |  | X
292 | 3 | 6 | 123 | 021A | 2SP | R/W | **** See note |  | X
293 | 3 | 6 | 124 | 020B | 3Ti | R/W | Time for segment. |  | X
294 | 3 | 6 | 125 | 021B | 3SP | R/W | **** See note |  | X
295 | 3 | 6 | 126 | 020C | 4Ti | R/W | Time for segment. |  | X
296 | 3 | 6 | 127 | 021C | 4SP | R/W | **** See note |  | X
297 | 3 | 6 | 128 | 020D | 5Ti | R/W | Time for segment. |  | X
298 | 3 | 6 | 129 | 021D | 5SP | R/W | **** See note |  | X
299 | 3 | 6 | 012A | 020E | 6Ti | R/W | Time for segment. |  | X
300 | 3 | 6 | 012B | 021E | 6SP | R/W | **** See note |  | X
301 | 3 | 6 | 012C | 020F | 7Ti | R/W | Time for segment. |  | X
302 | 3 | 6 | 012D | 021F | 7SP | R/W | **** See note |  | X
303 | 3 | 6 | 012E | 210 | 8Ti | R/W | Time for segment. |  | X
304 | 3 | 6 | 012F | 220 | 8SP | R/W | **** See note |  | X
305 | 3 | 6 | 130 | 211 | 9Ti | R/W | Time for segment. |  | X
306 | 3 | 6 | 131 | 221 | 9SP | R/W | **** See note |  | X
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
307 | 3 | 6 | 132 | 212 | 10Ti | R/W | Time for segment. |  | X
308 | 3 | 6 | 133 | 222 | 10SP | R/W | **** See note |  | X
309 | 3 | 6 | 134 | 213 | 11Ti | R/W | Time for segment. |  | X
310 | 3 | 6 | 135 | 223 | 11SP | R/W | **** See note |  | X
311 | 3 | 6 | 136 | 214 | 12Ti | R/W | Time for segment. |  | X
312 | 3 | 6 | 137 | 224 | 12SP | R/W | **** See note |  | X
313 | 3 | 6 | 138 | 215 | 13Ti | R/W | Time for segment. |  | X
314 | 3 | 6 | 139 | 225 | 13SP | R/W | **** See note |  | X
315 | 3 | 6 | 013A | 216 | 14Ti | R/W | Time for segment. |  | X
316 | 3 | 6 | 013B | 226 | 14SP | R/W | **** See note |  | X
317 | 3 | 6 | 013C | 217 | 15Ti | R/W | Time for segment. |  | X
318 | 3 | 6 | 013D | 227 | 15SP | R/W | **** See note |  | X
319 | 3 | 6 | 013E | 218 | 16Ti | R/W | Time for segment. |  | X
320 | 3 | 6 | 013F | 228 | 16SP | R/W | **** See note |  | X
321 | 3 | 6 | 140 | 024E | INPC | R/W | RD / WR in counts **** See note other paramters can reset to 0 | X | X | X
322 | 3 | 6 | 141 | 244 | LPBR | R/W | WR 1 = 1 second 0 = OFF | X | X | X
323 | 3 | 6 | 142 | 265 | SECR | R/W | 1 to 4 | X | X | X
324 | 3 | 6 | 143 | 247 | INPT | R/W | WR 1 = 0.1 minutes | X | X | X
325 | 3 | 6 | 144 | 245 | SENC | R/W | MUST pwr cycle to clr error SCZ | X | X | X
326 | 3 | 6 | 145 | 258 | SCAL | R/W | Read Only for INP 1 to 14 !! See note | X | X | X
327 | 3 | 6 | 146 | 259 | SCAH | R/W | Read Only for INP 1 to 14 !! See note | X | X | X
328 | 3 | 6 | 147 | 256 | SPL | R/W | !! See note | X | X | X
329 | 3 | 6 | 148 | 257 | SPH | R/W | !! See note | X | X | X
330 | 3 | 6 | 149 | 248 | S1OL | R/W | WRITE 50 for 50% | X | X | X
331 | 3 | 6 | 014A | 249 | S1OH | R/W | WRITE 50 for 50% | X | X | X
332 | 3 | 6 | 014B | 024A | S2OL | R/W | WRITE 50 for 50% | X | X
333 | 3 | 6 | 014C | 024B | S2OH | R/W | WRITE 50 for 50% | X | X
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
334 | 3 | 6 | 014D | 261 | POL | R/W | **** See note (Model PP45 only)
335 | 3 | 6 | 014E | 262 | POH | R/W | **** See note (Model PP45 only)
336 | 3 | 6 | 014F | 263 | RSCL | R/W
337 | 3 | 6 | 150 | 264 | RSCH | R/W
338 | 3 | 6 |  |  |  | R/W | Not implemented
339 | 3 | 6 |  |  |  | R/W | Not implemented
340 | 3 | 6 | 153 | 266 | MANUALSP1 | R/W | WRITE 500 for 50.0% (auto must be off) |  | X
341 | 3 | 6 | 154 | 267 | MANUALSP2 | R/W | WRITE 500 for 50.0% (auto must be off) |  | X
342 | 3 | 6 |  |  |  |  | illegal address for all
343 | 3 |  | 156 |  | PCTOUT | Read Only | READ 50 for 50% +#=SP1 -#=SP2 | X | X | X
344 | 3 | 6 | 157 | 269 | PRE1 | R/W | WRITE 50 for 50% |  | X
345 | 3 | 6 | 158 | 026A | PRE2 | R/W | WRITE 50 for 50% |  | X
346 | 3 | 6 | 159 | 026E | A1D | R/W | 1 to 9999 counts **** See note
347 | 3 | 6 | 015A | 026F | A2D | R/W | 1 to 9999 counts **** See note
348 | 3 | 6 | 015B | 026C | A1TD | R/W | 0 (off) to 8000 seconds
349 | 3 | 6 | 015C | 026D | A2TD | R/W | 0 (off) to 8000 seconds
350 | 3 | 6 |  |  | ALEVNT1_8 | R/W | bit=1=On bit=0=Off |  | X
|  |  | 120 | 024C |  | [High Byte] | bit7 Segment 1 A1 msb
|  |  | 120 | 024C |  |  | bit6 Segment 1 A2
|  |  | 122 | 024C |  |  | bit5 Segment 2 A1
|  |  | 122 | 024C |  |  | bit4 Segment 2 A2
|  |  | 124 | 024C |  |  | bit3 Segment 3 A1
|  |  | 124 | 024C |  |  | bit2 Segment 3 A2
|  |  | 126 | 024C |  |  | bit1 Segment 4 A1
|  |  | 126 | 024C |  |  | bit0 Segment 4 A2 lsb
|  |  | 128 | 024C |  | [Low Byte] | bit7 Segment 5 A1 msb
|  |  | 128 | 024C |  |  | bit6 Segment 5 A2
|  |  | 012A | 024C |  |  | bit5 Segment 6 A1
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
|  |  | 012A | 024C |  |  | bit4 Segment 6 A2
|  |  | 012C | 024C |  |  | bit3 Segment 7 A1
|  |  | 012C | 024C |  |  | bit2 Segment 7 A2
|  |  | 012E | 024C |  |  | bit1 Segment 8 A1
|  |  | 012E | 024C |  |  | bit0 Segment 8 A2 lsb
351 | 3 | 6 |  |  | ALEVNT9_16 | R/W | bit=1=On bit=0=Off |  | X
|  |  | 130 | 024D |  | [High Byte] | bit7 Segment 9 A1 msb
|  |  | 130 | 024D |  |  | bit6 Segment 9 A2
|  |  | 132 | 024D |  |  | bit5 Segment 10 A1
|  |  | 132 | 024D |  |  | bit4 Segment 10 A2
|  |  | 134 | 024D |  |  | bit3 Segment 11 A1
|  |  | 134 | 024D |  |  | bit2 Segment 11 A2
|  |  | 136 | 024D |  |  | bit1 Segment 12 A1
|  |  | 136 | 024D |  |  | bit0 Segment 12 A2 lsb
|  |  | 138 | 024D |  | [Low Byte] | bit7 Segment 13 A1 msb
|  |  | 138 | 024D |  |  | bit6 Segment 13 A2
|  |  | 013A | 024D |  |  | bit5 Segment 14 A1
|  |  | 013A | 024D |  |  | bit4 Segment 14 A2
|  |  | 013C | 024D |  |  | bit3 Segment 15 A1
|  |  | 013C | 024D |  |  | bit2 Segment 15 A2
|  |  | 013E | 024D |  |  | bit1 Segment 16 A1
|  |  | 013E | 024D |  |  | bit0 Segment 16 A2 lsb
352 | 3 |  |  |  | PCTOUT1 | Read Only | READ 50 for 50% SP1 Output | X | X | X
353 | 3 |  |  |  | PCTOUT2 | Read Only | READ 50 for 50% SP2 Output | X | X | ff # = On
620 |  | 6 |  | 026B | Write CAL8 | Write Only | Caution! Calibration Register
769 | 3 |  | 300 |  | OUT1 | Read Only | [A write to SP1D sets to ON/OFF] | X | X | X
00 or 01 = Time Proportioning
06 = Current or Voltage Output
08 = Pulse
16 = On/Off
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
770 | 3 | 6 | 301 | 229 | OUT1_TP | R/W | 1 to 80 | X | X | X
771 | 3 | 6 | 302 | 022D | OUT1_PUL | R/W | 1 to 7 | X | X | X
772 | 3 |  | 303 |  | OUT2 | Read Only | [A write to SP2D sets to ON/OFF] | X | X
00 or 01 = Time Proportioning
06 = Current or Voltage Output
08 = Pulse
16 = On/Off
773 | 3 | 6 | 304 | 022A | OUT2_TP | R/W | 1 to 80 | X | X
774 | 3 | 6 | 305 | 022E | OUT2_PUL | R/W | 1 to 7 | X | X
775 | 3 | 6 | 306 | 025D | TUNE or 1TUN | R/W | Rd 0,4,16,32,48,64 Wr 0 to 4 *see note | X | X | X
Rd 0 = Self Tune / Learn No
Rd 4 = Self Tune / Learn Yes
Rd 16 = PID
Rd 32 = Slow
Rd 48 = Normal
Rd 64 = Fast
Wr 0 = Self Tune
Wr 1 = PID
Wr 2 = Slow
Wr 3 = Normal
Wr 4 = Fast
776 | 3 | 6 | 307 | 025E | 2TUN | R/W | Refer to the 1TUN data above
777 | 3 | 6 | 308 | 025F | 3TUN | R/W | Refer to the 1TUN data above
778 | 3 | 6 | 309 | 260 | 4TUN | R/W | Refer to the 1TUN data above
779 | 3 | 6 | 030A | 023B | DFAC or 1DFAC | R/W | 0 to 7 | X | X | X
780 | 3 | 6 | 030B | 023C | 2DFAC | R/W | 0 to 7
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
781 | 3 | 6 | 030C | 023D | 3DFAC | R/W | 0 to 7
782 | 3 | 6 | 030D | 023E | 4DFAC | R/W | 0 to 7
783 | 3 | 6 | 030E | 415/414 | PID2 | R/W | 0 = Off # = On | X | X
784 | 3 | 6 | 030F | 417/416 | ARUP | R/W | 0 = Off # = On | X | X | X
785 | 3 | 6 | 310 | 40D/40C | PCTO | R/W | 0 = Off # = On | X | X
786 | 3 | 6 | 311 | 419/418 | PROG | R/W | 0 = Off # = On |  | X
787 | 3 | 6 | 312 | 41B/41A | PSET | R/W | 0 = Off # = On |  | X
788 | 3 | 6 | 313 | 41D/41C | STAT | R/W | 0 = Off # = On |  | X
789 | 3 | 6 | 314 | 420/421 | RUNHOLD | R/W | 0 = RUN # = HOLD |  | X
790 | 3 | 6 | 315 | 024F | PEND | R/W | 0 = Hold | X | X
1 = Ooff
2 = Loop
3 = SP1
791 | 3 | 6 | 316 | 246 | FILT | R/W | 0 to 99 | X | X | X
792 | 3 | 6 | 317 | 025A | INP | R/W | 1 = J-IC | X | X | X
2 = CA
3 = E
4 = T
5 = L
6 = N
7 = R-13
8 = S-10
9 = B
10 = C
[7 empty] | 11 = P392 |  |  | N.A.
[7 empty] | 12 = N120 |  |  | N.A.
[7 empty] | 13 = P385 |  |  | N.A.
14 = 1P38
15 = CURRENT
16 = VOLTAGE
[7 empty] | 17 = DIFFERENTIAL |  |  | N.A.
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
793 | 3 | 6 | 318 | 423/422 | OSUP | R/W | 0 = Off # = On | X | X | X
794 | 3 | 6 | 319 | 025B | UNIT | R/W | 0 = NONE | X | X | X
1 = DEG. F
2 = DEG. C
795 | 3 | 6 | 031A | 025C | DPT | R/W | 0 = 0 | X | X | X
1 = 0.0
2 = 0.00
3 = 0.000
[7 empty] | 4 = 0.0000 |  | N.A. | N.A.
796 | 3 | 6 | 031B | 252 | SP1SETUP | R/W | [High Byte] Not used | X | X | X
[6 empty] | [Low Byte] | bit 7 Not used
bit6 Not used
bit5 S1iH 1 = On 0 = Off
bit4 S1Pi 1 = On 0 = Off
bit3 S1rE 1 = On/Off 0 = Hold
bit2 S1LP 1 = Oon 0 = Ooff
bit1 S1St 1 = dir 0 = rE
bit0 Not used
797 | 3 | 6 | 031C | 253 | SP2SETUP | R/W | [High byte] Not used | X | X
[6 empty] | [Low Byte] | bit 7 Not used
bit6 Not used
bit5 S2iH 1 = On 0 = Off
bit4 S2Pi 1 = On 0 = Off
bit3 S2rE 1 = On/Off 0 = Hold
bit2 S2LP 1 = Oon 0 = OoFF
bit1 S2St 1 = dir 0 = rE
bit0 S2t 1 = AbS 0 = dE
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
798 | 3 | 6 | 031D | 250 | AL1 | R/W | 0 = Off | X | X
1 = Lo
2 = Hi
3 = HiLo
[7 empty] | 4 = Evnt | N.A.
799 | 3 | 6 | 031E | 254 | AL1SETUP | R/W | [High byte] Not used | X | X
[6 empty] | [Low Byte] | bit7 Not used
bit6 A1Lb 1 = On 0 = Off
bit5 A1iH 1 = On 0 = Off
bit4 A1Pi 1 = On 0 = Off
bit3 A1rE 1 = OnOF 0 = Hold
bit2 A1LP 1 = Oon 0 = OoFF
bit1 A1St 1 = OPEn 0 = CLOS
bit0 A1t 1 = AbS 0 = dE
800 | 3 | 6 | 031F | 251 | AL2 | R/W | 0 = Off
1 = Lo
2 = Hi
3 = HiLo
4 = Evnt
801 | 3 | 6 | 320 | 255 | AL2SETUP | R/W | [High byte] Not used
[6 empty] | [Low Byte] | bit7 Not used
bit6 A2Lb 1 = On 0 = Off
bit5 A2iH 1 = On 0 = Off
bit4 A2Pi 1 = On 0 = Off
bit3 A2rE 1 = OnOF 0 = Hold
bit2 A2LP 1 = Oon 0 = OoFF
bit1 A2St 1 = OPEn 0 = CLOS
bit0 A2t 1 = AbS 0 = dE
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
802 | 3 | 6 | 321 | 425/424 | POSR | R/W | 0 = InP # = SPt
803 | 3 | 6 | 322 | 410 | SP | R/W | 0 = 1SP1
|  |  |  | 411 |  |  | 1 = 2SP1
|  |  |  | 412 |  |  | 2 = 3SP1
|  |  |  | 413 |  |  | 3 = 4SP1
804 | 3 | 6 | 323 | 40F/40E | SPSA | R/W | 0 = Int # = rE
805 | 3 | 6 | 324 | 401/400 | LOrE | R/W | 0 = LOC # = rE | X | X | X
806 | 3 | 6 | 325 |  | nAt | Read Only | 0 to 99 | X | X | X
807 | 3 | 6 | 326 | 427/426 | RSPT | R/W | 0 = Off # = On
808 | 3 | 6 |  |  | ISEN | R/W | Not functional on ALL
809 | 3 | 6 | 328 | 432/433 | SP1O | R/W | 0 = OutA # = OutB | X | X | X
810 | 3 | 6 | 329 | 268 | InPB | R/W | 0 = FAIL |  | X
1 = AVE
2 = PRE
811 | 3 | 6 | 032A | 434/435 | APCT | R/W | 0 = Real # = Adj |  | X
812 | 3 | 6 | 032B | 436 | SP2_CONT_AL | R/W | 0 = AL # = Cont |  | X
813 | 3 | 6 | 032C | 440 | STOR | R/W | 0 =No (WR to RAM) # =Yes (EEPROM) | X | X | X
1025 |  | 6 |  |  | LOrE_rE | Write Only | **note | X | X | X
1026 |  | 6 |  |  | LOrE_LOC | Write Only | **note | X | X | X
1027 |  | 6 |  |  | RESET_ENTR | Write Only | Not functional on non-M300
1028 |  | 6 |  | 403 | ACK_AL1 | Write Only |  | X | X
1029 |  | 6 |  | 404 | ACK_AL2 | Write Only
1030 |  | 6 |  |  | ACK_AL12 | Write Only
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
1031 |  | 6 |  | 406 | ACK_SP1 | Write Only |  | X | X | X
1032 |  | 6 |  | 407 | ACK_SP2 | Write Only |  | X | X
1033 |  | 6 |  | 408 | AUTO_ON | Write Only |  |  | X
1034 |  | 6 |  | 409 | AUTO_OFF | Write Only |  |  | X
1035 |  | 6 |  | 040A | RESET_PEA | Write Only |  | X | X | X
1036 |  | 6 |  | 040B | RESET_VAL | Write Only |  | X | X | X
1037 |  | 6 |  |  | PCTO_ON | Write Only | **note | X | X
1038 |  | 6 |  |  | PCTO_OFF | Write Only | **note | X | X
1039 |  | 6 |  |  | SPSA_RE | Write Only | **note
1040 |  | 6 |  |  | SPSA_INT | Write Only | **note
1041 |  | 6 |  |  | SP_1SP1 | Write Only | **note
1042 |  | 6 |  |  | SP_2SP1 | Write Only | **note
1043 |  | 6 |  |  | SP_3SP1 | Write Only | **note
1044 |  | 6 |  |  | SP_4SP1 | Write Only | **note
1045 |  | 6 |  |  | PID2_ON | Write Only | **note | X | X
1046 |  | 6 |  |  | PID2_OFF | Write Only | **note | X | X
1047 |  | 6 |  |  | ARUP_ON | Write Only | **note | X | X | X
1048 |  | 6 |  |  | ARUP_OFF | Write Only | **note | X | X | X
1049 |  | 6 |  |  | PROG_ON | Write Only | **note |  | X
1050 |  | 6 |  |  | PROG_OFF | Write Only | **note |  | X
1051 |  | 6 |  |  | PSET_ON | Write Only | **note |  | X
1052 |  | 6 |  |  | PSET_OFF | Write Only | **note |  | X
1053 |  | 6 |  |  | STAT_ON | Write Only | **note |  | X
1054 |  | 6 |  |  | STAT_OFF | Write Only | **note |  | X
1055 |  | 6 |  | 041E | TBAS_1 | Write Only |  |  | X
1056 |  | 6 |  | 041F | TBAS_60 | Write Only |  |  | X
1057 |  | 6 |  |  | RunHold_RUN | Write Only | **note |  | X
1058 |  | 6 |  |  | RunHold_HOLD | Write Only | **note |  | X
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
1059 |  | 6 |  |  | OSUP_ON | Write Only | **note | X | X | X
1060 |  | 6 |  |  | OSUP_OFF | Write Only | **note | X | X | X
1061 |  | 6 |  |  | POSR_SPT | Write Only | **note
1062 |  | 6 |  |  | POSR_INP | Write Only | **note
1063 |  | 6 |  |  | RSPT_ON | Write Only | **note
1064 |  | 6 |  |  | RSPT_OFF | Write Only | **note
1065 |  | 6 |  |  | ISEN_ON | Write Only | NON-FUNCTIONAL ALL
1066 |  | 6 |  |  | ISEN_OFF | Write Only | NON-FUNCTIONAL ALL
1067 |  | 6 |  | 042A | 1LEARN_YES | Write Only |  | X | X | X
1068 |  | 6 |  | 042B | 1LEARN_NO | Write Only |  | X | X | X
1069 |  | 6 |  | 042C | 2LEARN_YES | Write Only
1070 |  | 6 |  | 042D | 2LEARN_NO | Write Only
1071 |  | 6 |  | 042E | 3LEARN_YES | Write Only
1072 |  | 6 |  | 042F | 3LEARN_NO | Write Only
1073 |  | 6 |  | 430 | 4LEARN_YES | Write Only
1074 |  | 6 |  | 431 | 4LEARN_NO | Write Only
1075 |  | 6 |  |  | SP1O_OUTA | Write Only | **note | X | X | X
1076 |  | 6 |  |  | SP1O_OUTB | Write Only | **note | X | X | X
1077 |  | 6 |  |  | APCT_REAL | Write Only | pcto is real% of control output **note |  | X
1078 |  | 6 |  |  | APCT_ADJ | Write Only | is adj.% based on S1OL/S1OH **note |  | X
1079 |  | 6 |  |  | SP2_CONT | Write Only | **note |  | X
1080 |  | 6 |  |  | SP2_AL | Write Only | **note |  | X
1089 |  | 6 |  | 440 | STOR_YES | Write Only | write commands are to EEPROM **note | X | X | X
1090 |  | 6 |  | 441 | STOR_NO | Write Only | write commands are to RAM **note | X | X | X
1091 |  | 6 |  | 442 | STORE_ALL | Write Only | copy RAM data to EEPROM | X | X | X
1793 | 3 |  | 700 |  | COMM VER # | Read Only | Week/Year in HEX. (ex. 15 00) | X | X | X
1794 | 3 |  | 701 |  | MODEL I.D. # | Read Only | ASCII Char. (ex. 53 5A = S Z) | X | X | X
1795 | 3 |  | 702 |  | SOFT VER # | Read Only | Week/Year in HEX. (ex. 09 00) | X | X | X
MODBUS |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
1841 | 3 |  | 730 |  | Read CAL 2 | Read Only | Calibration Register
1842 | 3 |  | 731 |  | Read CAL 3 | Read Only | Calibration Register
1843 | 3 |  | 732 |  | Read CAL 4 | Read Only | Calibration Register
1844 | 3 |  | 733 |  | Read CAL 5 or 6 | Read Only | Calibration Register
1845 | 3 |  | 734 |  | Read CAL 7 | Read Only | Calibration Register
1846 | 3 |  | 735 |  | Read CAL 8 | Read Only | Calibration Register
1847 |  | 6 |  | 736 | Write CAL 5 | Write Only | Caution! Calibration Register
1848 |  | 6 |  | 737 | Write CAL 6 | Write Only | Caution! Calibration Register
1849 |  | 6 |  | 738 | FACT DFLT | Write Only | Unit reboots, gives no response
[6 empty] | *NOTE: | To write the low nibble (learn yes/no)
you must use the corresponding
write only command. For example to
write the 1TUN yes/no bit use the
1067 1learnyes or 1068 1learnno cmds.
[6 empty] | **NOTE: | Also exists as a R/W register
[6 empty] | ***NOTE: | Writes to registers above #1024 do not
recognize specific values, any data value
can be used.
[6 empty] | ****NOTE: | Apply the DPT (0x031A) setting
to the register value.
[6 empty] | !! NOTE: | For INP 15,16,17 apply the DPT setting
[6 empty] | && NOTE: | Indicates special difference from Modbus command
BIT |  |  | Love Link
REGISTER | Read | Write | Read | Write | Parameter | Operation | Notes | CN 78000 | CN 74000 | CN 79000
1 | 4 |  |  |  | Reserved | Read Only
2 | 4 |  |  |  | CHEC LORE | Read Only | No Activity Timer timeout 0=FALSE 1=T | X | X | X
3 | 4 |  |  |  | AL2 | Read Only | Alarm 2 output state 0=OFF 1=ON |  | X
4 | 4 |  |  |  | AL1 | Read Only | Alarm 1 output state 0=OFF 1 | X | X
5 | 4 |  |  |  | ERROR PRESENT | Read Only | 0=FALSE 1=TRUE | X | X | X
6 | 4 |  |  |  | LOC/REM | Read Only | 0=LOCAL 1=REMOTE | X | X | X
7 | 4 |  |  |  | SENC BAD | Read Only | RATE OF CHANGE > LIMIT 0=FALSE 1=T | X | X | X
8 | 4 |  |  |  | AREA | Read Only | AMBIENT BEYOND SPEC 0=FALSE 1=T | X | X | X
9 | 4 |  |  |  | OPEN INP | Read Only | 0=FALSE 1=TRUE | X | X | X
10 | 4 |  |  |  | BAD INP | Read Only | 0=FALSE 1=TRUE | X | X | X
11 | 4 |  |  |  | UFL | Read Only | 0=FALSE 1=TRUE | X | X | X
12 | 4 |  |  |  | OFL | Read Only | 0=FALSE 1=TRUE | X | X | X
13 | 4 |  |  |  | CHEC CAL | Read Only | 0=FALSE 1=TRUE | X | X | X
14 | 4 |  |  |  | FAIL TEST | Read Only | 0=FALSE 1=TRUE | X | X | X
15 | 4 |  |  |  | LOOP BAD | Read Only | 0=FALSE 1=TRUE | X | X | X
16 | 4 |  |  |  | AUTO/MANUAL | Read Only | 0=AUTO 1=MANUAL | X | X | X